BME 140 W24

Lab #1 Report

Lab: B1-2

Equipment Familiarization and Measurement Theory

Brian Ngo

Due: 1/25/24

Results and Discussion

In section A, we measured the voltage output of the GPE-2323 power supply with a digital multimeter (DMM). The expected voltage is 10V and what we recorded is 9.94 ± 0.01V. With a fluctuating value, which approaches near 10V, it seemed everything was calibrated correctly, and our digital multimeter was working without problems. Next, we used an oscilloscope getting the value 9.730 ± 0.015V. Even closer to our expected voltage, the oscilloscope was calibrated and working as intended as well.

In Section B, my lab partner and I chose a one kOhm and a 7.5 MOhm resistor to calculate the accuracy of Rnominal and RDMM. Here are our results:

|  | 100 Ohm – 10 kOhm | 100 kOhm - 10 MOhm |
| --- | --- | --- |
| Colors | Black Brown Red | Violet Green Green |
| Rnominal | 1 kOhm | 7.5 MOhm |
| Rtolerance | Gold (5%) | Gold (5%) |
| RDMM | 0.992 ± 0.001 kOhm | 7.70 ± 0.01 MOhm |

To calculate the accuracy of our two resistors. We use this formula.

|Expected Value-  Calculated Value|Expected Value*100

In this case:

Black Brown Red:  |1-0.992|1*100 = 0.8%

Violet Green Green: |7.5-7.70|7.5*100 = 2.6%

Both of our resistors were within the accuracy tolerance of 5%.

In Section C, we are asked to predict the current through the “kOhm” resistor using VPS,DMM, and RDMM.

I±∆I=fV,R±[∂f∂V∆V+∂f∂R∆R]

We can use our measured voltage obtained in section A and since we are looking at only the “kOhm” resistor, we can use our RDMM value for our “Brown Brown Red” resistor to find current since the other is in “MOhm”.

I=VR=9.940.992=0.010 mA

And to find our change in current.

∆I= -9.940.992*1032*0.001*103+10.992*103*0.01=2.018*10-5

So, I=0.010±2.015*10-5A= 0.010±2.015*10-2mA

In Section D, we were tasked to measure the frequency of an SDG-830 function generator at 100 Hz. Wiring it to the oscilloscope, we found the value to display 100 ± 0.27, so the machine was calibrated and running as intended. To maximize the resolution, adjusting the time base settings to the minimum and the vertical voltage scale to the maximum allows for more data points within each waveform cycle and ensures small variations in voltage are visible.

In Section E, we are asked to observe the sine wave from the same function generator in section D, setting the sine wave output to a peak-to-peak amplitude of 10.000V (with an uncertainty of ± 0.001). We can use this value to calculate the root-mean-square voltage. Since it is peak-to-peak, we would divide 10 in half to represent an amplitude from the x-axis.

Amplitude√2=5√2=3.54V.

Additionally, measuring the DC and root-mean-square voltages using the DMM, we get the following:

Vrms = 3.554 ± 0.001V

and

VDC,DMM = -10.7 ± 0.1 mV

Note: rms represents the AC voltage value

In section F, we are asked to set up a voltage divider and measure the voltage (V) across the bottom resistor using the DMM and the oscilloscope with the following resistors: 1kOhm, 110kOhm, and 10MOhm. Additionally, an impedance of 50 Ohm and 1MOhm were set on the oscilloscope and recorded for each resistor. Here are our results:

| Resistance | Calculated | VDMM | VOsc (50 Ohms) | VOsc (1 MOhm) |
| --- | --- | --- | --- | --- |
| 1 kOhms | 7.5V | 7.48 ± 0.01 V | 704.1 ± 1.0 mV | 7.433 ± 0.010 V |
| 110 kOhms | 7.5V | 7.47 ± 0.01V | 23.50 ±7.00 mV | 7.078 ± 0.003 V |
| 10 MOhm | 7.5V | 5.108 ± 0.001V | 16.00 ±1.5 mV | 1.131 ± 0.003 V |

The calculated voltage value is 7.5V throughout. This is because it is 15V going through 2 resistors in series. The discrepancy between the three resistors in our calculated and digital multimeter (DMM) voltages is not far off except for 10 MOhm by a difference of 2.392V. Comparing the data between DMM and the oscilloscopes, the measurements vary. The 50 Ohm resistance was a lot less (1 volt = 1000 mV), so we were getting values less than 1V, not remotely close to both the calculated and DMM values. 1 MOhm measurements, however, were a lot closer (except for our 10 MOhm recordings). This makes sense because the higher input impedance (1MOhm) imposes less load on the circuit under test, reducing the input on the voltage being measured.

The DMM should have the greater internal resistance because an ideal voltmeter should not be able to draw current from the circuit, hence the resistance of the voltmeter must be infinity. To calculate the internal resistance of the DMM, we would use any of the results from our DMM readings and the current to calculate internal resistance (Rin) which would be the internal impedance/resistance of the device. The result is around 10 MOhm. It’s important to note that the more accurate voltage measuring device will be the one with the least attenuation, so that would be the DMM.

If we want to find the required minimum input resistance of the measuring instrument if R = 10 MOhms in Figure 2 and we want to measure voltage with an error less than 1%, we will just mathematically solve it with a given equation and a design rule of thumb. Additionally, since circuit A drives circuit B, to not attenuate the signal, we must have Zout, a << Zin,b

The following equation is used:

Vout, a=Zin, bZin, b+Zout,a*Vs

We know Zin,b = 20 MOhm and Zout,a = RT = 12R (resistors are parallel). Additionally, R << 40 MOhm. Using the design rule of thumb, we get R=4 MOhm. So, solving for Zout, a, we get 0.91 Vs, which is less than 1% error. Furthermore, we were also tasked with measuring the resistance of a 7.5 MOhm resistor with and without touching the resistors. Here are our results:

With: 130 ± 10 kOhm

Without: 7.65 ± 0.01 MOhm

Conclusion

In this laboratory experiment, we delved into the fundamental aspects of basic circuits and instruments, focusing on components such as resistors, breadboards, digital multimeters (DMM), function generators, oscilloscopes, and dual-output power supplies. The understanding of these components is crucial for constructing and analyzing electronic circuits.

We began by deciphering resistor values through their color codes and gaining insights into their tolerance levels and resistance values. The utilization of digital multimeters allowed us to measure DC and AC voltages, currents, and resistance values accurately. The function generator, with its ability to produce diverse waveforms and modulation, offered a versatile tool for signal generation. The oscilloscope, an advanced DPO, was employed to visualize signals with precision. We familiarized ourselves with its controls, including auto scale, voltage, and time divisions, triggering sources, and measurement functions. These tools collectively provided a comprehensive understanding of signal behavior and characteristics. Additionally, the dual output power supply was introduced as a reliable source of constant voltage and current, essential for various circuit applications.